## Slide 1
Maciej Budzanowski, Renata Kopeć, Anna Bubak, Mariusz Kłosowski, Barbara Marczewska, Pawel Olko, Anna Sas-Bieniarz
Progress in thermoluminescence dosimetry for individual and environmental monitoring
Institute of Nuclear Physics, Kraków, Poland

## Slide 2
Individual and environmental dosimeter essential for operation of Nuclear Power Plant
After Chernobyl it was impossible to assess individual doses due to:
lack of individual dosemeters
too low dose range of dosimetric films

## Slide 3
TLDs have very good dosimetric properties
Thermoluminescent dosimeters (TLDs) broadly used in individual and environmental monitoring (also at NPP)
Some disadvantages:
- 2nd readout not possible
- no information on static exposure
- limited dose range in accident
How to improve thermoluminescence dosimetry?
TLD dosimetry at LADIS Dosimetric Service at IFJ PAN

## Slide 4
To develop procedure of dose verification (second readout of TLD)
2) To get information on static exposure
3) ) To increase the dose range to 10 Gy
The aim of the work

## Slide 5
Introduction to thermoluminescence and thermoluminescent dosimetry

## Slide 6
Maria Curie-Skłodowska
1904 – observed thermoluminescence in natural fluorite CaF2
„....the radium source restores the thermoluminescence of fluorytu”

## Slide 7
1960 First LiF monocrystals for dosimetry
1974 Patent of T. Niewiadomski LiF:Mg,Ti (MTS-N)
1978 Licence for Studsvik for MTS-N
1983 Thin layer CaSO4:Dy for radon dosimetry
1985 High sensitive MCP-N (LiF:Mg,Cu,P)
1995 Thin layer MCP-Ns
2005 MTT (LiF:Mg,Ti) for space measurements kosmosie
2005 Detektory powierzchniowe 2-D do pomiarów klinicznych
Tadeusz Niewiadomski, Institute of Nuclear Physics , Kraków, Poland (1920-1996)

## Slide 8
Model of luminescence - excitation
Yukihara, SSD 15, 2007

## Slide 9
Model of luminescence - stimulation
Yukihara, SSD 15, 2007
Heat
Luminescence

## Slide 10
Ionisation and trapping holes and electrons in trapping centers
Thermoluminescence – 2 stage process
Notes: You know the priciple of thermoluminescence
Ionising radiation, interacting with TL detectors, produces ionisa

## Slide 11
Ionisation and trapping holes and electrons in trapping centers
Thermoluminescence – 2 stage process
Photomultiplier
Notes: You know the priciple of thermoluminescence
Ionising radiation, interacting with TL detectors, produces ionisa

## Slide 12
What do we need ? TL detectors
LiF:Mg,Ti LiF:Mg,Cu,P CaSO4:Dy

## Slide 13
CCD
PM
High –sensitive
CCD camera
What do we need ? TL readers

## Slide 14
What do we need ? TL readers
Manual RA-94 TLD reader (Mikrolab, Kraków)
Experimental 2D TLD reader with CCD camera, IFJ PAN Kraków
Notes: You know the priciple of thermoluminescence
Ionising radiation, interacting with TL detectors, produces ionisa

## Slide 15
Result of readout : TL glow curve

## Slide 16
Second readout of TL detectors for re-estimation of dose

## Slide 17
TL light
conduction band
valence band
radiation
conduction band
valence band
Phenomenon of luminescence

## Slide 18
MTS-N ( LiF:Mg,Ti) detectors, 50 mGy
Normal readout of TL detector (LiF:Mg,Ti)

## Slide 19
conducting band
valence band
conducting band
valence band
TL light
UV radiation
UV light on irradiated TL detector

## Slide 20
Readout
UV exposure
2nd readout
Irradiation (g)
The method of work

## Slide 21
2nd readout after UV irradiation of detector (LiF:Mg,Ti)

## Slide 22
MTS-N ( LiF:Mg,Ti) detectors glow curves after the 20 mGy irradition.
1st readout – normal 2nd readout - after UV irradiation

## Slide 23
Detector type
UV wavelenght
  - 254 nm
  - 302 nm
  - 365 nm
Exposure time
Background correction
Optmalization of 2nd readout: UV exposure

## Slide 24
Optmalization of 2nd readout: UV wavelenght
MTS-N ( LiF:Mg,Ti) detectors glow curves after the 20 mGy irradition.
2nd readout - after UV irradiation 254 nm optimal

## Slide 25
Optmalization 2nd readou: time of UV exposure
Optimal irradiation time: 1.5 hours

## Slide 26
Optmalization of 2nd readout : temperaturę of UV exposure
Optimal irradiation temperature : 80 C

## Slide 27
Optmalization of 2nd readout: dose response

## Slide 28
10 mGy
50 mGy
Optmalization of 2nd readout : background corection

## Slide 29
Optimization of 2nd readout: summary
The re-estimation of dose can be performed with precision better than 10% for doses exceeding 10 mSv.
The method works for LiF:Mg,Ti detectors but not for high sensitive MCP-N (LiF;Mg,Cu,P).
The method is ready for use. It will be implemented at IFJ PAN - LADIS laboratory in Kraków, as a first in the world.

## Slide 30
Evaluation of doses up to 10 Gy

## Slide 31
RE-2000 TLD Reader
The problem
In individual dosimetry readers optimized for low doses
If D > 1 Sv appears then high probability of saturation of individual dosemeter
The aim of the work
To investigate, calibrate and test the system from low doses to 10 Gy

## Slide 32
TL glow curves after high doses of gamma-rays
LiF:Mg,Ti
LiF:Mg,Cu,P
LiF:Mg,Ti
LiF:Mg,Cu,P
10 Gy
1 Gy
10 Gy
1 Gy

## Slide 33
Przykład zastosowania dodatkowego filtra podczas odczytu dawek 10 Gy
Dla detektorów MTS-N (góra) i MCP-N (dół)
LiF:Mg,Cu,P
LiF:Mg,Ti
10 Gy
10 Gy
Installation of grey filter to reduce the emission intensity

## Slide 34
LiF:Mg,Cu,P
LiF:Mg,Ti
10 Gy
10 Gy
Readout with additional filter

## Slide 35
After exceeding the max. channel count
STOP
INTRODUCE FILTER
DETECTORS 2,3,4 measured with the filter
Method

## Slide 36
Results for MTS-N (LiF;Mg,Ti) detectors
supralinearity

## Slide 37
Results for MCP-N (LiF:Mg,Cu,P) detectors

## Slide 38
Identification of static exposure

## Slide 39
Films
(IMP Łodzi, Poland
TLDs
TLD (IFJ PAN Krakow)
TLD curve
Film exposed to static and dynamic exposure.
What is static and dynamic exposure?

## Slide 40
Method

## Slide 41
D= 20 mSv; static exposure
under Oo, 30o, 45o, 60o.
Dose 5 mSv, 10 mSv, 15 mSv i 20 mSv; dynamic exposure
40 kVp X-rays
Results for X-rays

## Slide 42
The problem
In nuclear industry exposure takes place at higher photon energies
In LADIS dosimetric service the highest exposures measured regularly for Cs-137 sources.
For 661 keV thicker filters and more sensitive cameras needed
Filter thickness
Photon Energy

## Slide 43
New CCD camera
MCP-N detectors covered with 1 mm, 2 mm and 4 mm Pb filters
Preliminary setup

## Slide 44
Pb thickness | Detector No. | TL without Pb | TL under Pb | Difference
1 mm | R88-1 | 266 | 258 | 8
 | R88-2 | 267 | 261 | 6
 | R88-3 | 268 | 260 | 8
 | R89-1 | 274 | 266 | 8
 | R89-2 | 273 | 265 | 8
 | R89-3 | 273 | 266 | 7
2 mm | R90-1 | 253 | 244 | 9
 | R90-2 | 248 | 239 | 9
 | R90-3 | 247 | 238 | 9
 | R94-1 | 255 | 245 | 10
 | R94-2 | 257 | 248 | 9
 | R94-3 | 261 | 252 | 9
4 mm | R95-1 | 238 | 219 | 19
 | R95-2 | 232 | 217 | 15
 | R95-3 | 239 | 223 | 16
 | R96-1 | 244 | 223 | 21
 | R96-2 | 232 | 214 | 18
 | R96-3 | 239 | 222 | 17
Preliminary results

## Slide 45
Method of analysis – histogram of pixel intensity

## Slide 46
Summary and conclusions
New method for re-estimation of dose for thermoluminescence detectors LiF:Mg,Ti js been developed. Second readout after 1.5 h UV irradiation allows to determine doses above 5 mSv.
The dose range for individual dosemeter has been extended to 10 Gy, allowing for using it in accidental exposures
The preliminary tests demenstrate that it is possible to distinguish statiic exposure of TLD detectors for 661 keV gamma-rays

## Slide 47
This research was kindly supported from the project
DEVELOPMENT OF NUCLEAR SAFETY AND RADIOLOGICAL PROTECTION METHODS
FOR THE NUCLEAR POWER ENGINEERING’S CURRENT AND FUTURE NEEDS
from
Narodowe Centrum Badan i Rozwoju, NCBiR